## Slide 1
Splash
Login
View Sequence
Run Sequence
Edit Sequence
Copy Sequence
Delete Sequence
Abort
Home
Manual
Run
Select Sequence
Create Sequence
Delete Operation
Prompt
Screen
Key
Error
Validation Warning
Abort
Error
Save
Complete

## Slide 2
Splash
ELIXYS
Automated Radiochemical Synthesis Platform
Designed and developed by the van Dam Lab
California Nanosystems Institute
University of California Los Angeles

## Slide 3
Login
Enter the server URL and your credentials:
Server
192.168.0.108
Username
hsimpson
Password
**********
Connect
Exit
This screen will only be needed when running the installed client (AIR)

## Slide 4
Connecting
Enter the server URL and your credentials:
Server
192.168.0.108
Username
hsimpson
Password
**********
Connect
Exit
Please wait, connecting…
Exit
This screen will only be needed when running the installed client (AIR)

## Slide 5
Home
What would you like to do?
Create, View or Run a Sequence
Operate the System Manually
Observe the Current Run
Most users will see this home page after logging in. Exception: The user who is currently running the Elixys system will go directly to the run screen.
This option will only appear if another user is currently
running the Elixys system.

## Slide 6
Select Sequence (Saved Sequences)
Back
Saved Sequences
Manual Runs
Create
View
Run
Copy
Delete
FAC
FAC (high temp)
FAC (long reaction time)
FDG
FLT

## Slide 7
Select Sequence (Manual Runs)
Back
Saved Sequences
Manual Runs
Create
View
Run
Copy
04-25-2011 6:53 am FAC
04-22-2011 12:06 pm FLT
04-17-2011 8:00 am FDG
04-14-2011 3:14 pm FAC
04-11-2011 3:14 am FAC
04-05-2011 12:34 pm Initial machine test
Copies of manual runs
are always saved under
Saved Sequences.

## Slide 8
Create Sequence
Back
Saved Sequences
Manual Runs
Create
View
Run
Copy
Delete
FAC
FAC (high temp)
FAC (long reaction time)
FDG
FLT
Enter the name of the new sequence:
Create
FAC (my new sequence)
Cancel

## Slide 9
Copy Sequence
Back
Saved Sequences
Manual Runs
Create
View
Run
Copy
Delete
FAC
FAC (high temp)
FAC (long reaction time)
FDG
FLT
Enter the name of the new sequence:
Copy
FAC (my duplicate sequence)
Cancel

## Slide 10
Delete Sequence
Back
Saved Sequences
Manual Runs
Create
View
Run
Copy
Delete
FAC
FAC (high temp)
FAC (long reaction time)
FDG
FLT
Are you sure that you want to permanently delete sequence
“FAC (high temp)”?
Delete
Cancel

## Slide 11
View Sequence (Cassette Configuration)
FAC (high temp)
Back
Edit
Run
1. Add F-18
2. Evaporate
4. Transfer
5. Elute
Previous
Next
Cassette 1
Cassette 2
Cassette 3
3. React
Position | Name | Description
1 | F-18 | [18F]F-, 10 mg Kryptofix (K222) and 1.0 mg potassium carbonate (K2CO3) in acetonitrile (MeCN)
2 | MeCN1 | 1 mL acetonitrile (MeCN)
3 | MeCN2 | 1 mL acetonitrile (MeCN)
4 | Precursor1 | 10 mg tribenzoyl pentose triflate in 0.5 mL acetonitrile (MeCN)
5 | EtOAc1 | 1 mL ethyl acetate
6 | EtOAc2 | 1 mL ethyl acetate
7 | EtOAc3 | 1 mL ethyl acetate
8 | EtOAc4 | 1 mL ethyl acetate
A | Column1 | Silica gel cartridge
Click/tap an operation in the left panel to change the contents of the right panel.

## Slide 12
View Sequence (Unit Operation)
FAC (high temp)
React unit 1, position 1 at 160 ⁰C for 15 minutes and cool to 35 ⁰C. Stir at 5000 rpm.
3. React
Flick up and down in this panel to navigate using multitouch.
1. Add F-18
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cassette 3
Cassette 1
Back
Edit
Run
Previous
Next

## Slide 13
Edit Sequence (Cassette Configuration)
FAC (high temp)
Back
Run
Previous
Next
Cassette 1
1. Add F-18
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cassette 3
3. React
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Use this cassette
Name:
Description:
F-18
[18F]F-, 10 mg Kryptofix (K222) and 1.0 mg potassium carbonate (K2CO3) in acetonitrile (MeCN)
1
Name:
Description:
MeCN1
1 mL acetonitrile (MeCN)
2
Name:
MeCN2
3
Tap/click to enable or disable this cassette.
A unique name is required, description is optional.

## Slide 14
Edit Sequence (Add/Insert)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add F-18
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cassette 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
160.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
Drag operation from toolbox and drop in sequence to add/insert.

## Slide 15
Edit Sequence (Reorder/Delete)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add F-18
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cassette 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
160.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
Drag operation up or down to reorder.
Drag to trash can to delete.

## Slide 16
Edit Sequence (Set Parameters)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add F-22
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cassette 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
160.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
Click/tap an operation in the middle panel to change the contents of the right panel. Set the operation parameters in the right panel.

## Slide 17
Edit Sequence (Delete)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cartridge 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
160.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
Are you sure that you want to delete operation 3 (React)?
Delete
Cancel

## Slide 18
Edit Sequence (Validation 1)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add F-18
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cassette 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
4000.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
!
!
Icons used to visually indicate validation errors. Tap icon to view the error message.

## Slide 19
Edit Sequence (Validation 2)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cartridge 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
4000.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
!
!
Validation error: The reaction temperature must be below 473 K.
OK

## Slide 20
Edit Sequence (Validation 3)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add F-18
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cassette 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
4000.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
!
!
Validation errors will cause a warning when Back is clicked or an error when Run is clicked.

## Slide 21
Edit Sequence (Validation 4)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cartridge 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
4000.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
!
!
The sequence “FAC (high temp)” has validation errors. Are you sure you want to go back to the sequence selection screen?
Back
Cancel

## Slide 22
Edit Sequence (Validation 5)
FAC (high temp)
Back
Run
Previous
Next
3. React
1. Add
2. Evaporate
4. Transfer
5. Elute
Cassette 2
Cartridge 3
Cassette 1
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Trash
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
4000.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
!
!
The sequence “FAC (high temp)” cannot be run because it contains validation errors.
OK

## Slide 23
Run Sequence
FAC (high temp)
Abort
1. Add F-18
2. Evaporate
3. React
4. Transfer
5. Elute
6. Install
7. Comment
8. Radiation
Advances automatically with the run (scrollable but not selectable).
Reacting unit 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm
Unit 1 Video

## Slide 24
Reacting unit 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm
Run Sequence (Abort)
FAC (high temp)
Abort
1. Add F-18
2. Evaporate
3. React
4. Transfer
5. Elute
6. Install
7. Comment
8. Radiation
Unit 1 Video
Are you sure you want to abort the sequence run? This operation cannot be undone.
Abort
Cancel

## Slide 25
Reacting unit 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm
Run Sequence (Error)
FAC (high temp)
Abort
1. Add F-18
2. Evaporate
3. React
4. Transfer
5. Elute
6. Install
7. Comment
8. Radiation
Unit 1 Video
Fatal error: The communication channel to the Elixys system has been lost and could not be restored. The run will now abort.
OK

## Slide 26
Reacting unit 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm
Run Sequence (Complete)
FAC (high temp)
Abort
1. Add F-18
2. Evaporate
3. React
4. Transfer
5. Elute
6. Install
7. Comment
8. Radiation
Unit 1 Video
Sequence run “FAC (high temp)” completed successfully.
OK

## Slide 27
Manual Run (Select Operation)
Sequence
Run Complete
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Addition of acetonitrile to unit 1 completed successfully.
Please select the next operation from the toolbox.
Move
Prompt
1. Add F-18
2. Evaporate
3. React
Cassette 2
Cassette 3
Cassette 1
Click/tap to select the next unit operation
Read only (scrollable but not selectable)
3. Add MeCN

## Slide 28
Manual Run (Configure Operation)
Sequence
React
Back
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Unit:
1
15
Duration:
Reaction temp:
Final temp:
Stir speed:
4000.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1
Configure Reaction Parameters
Move
Prompt
1. Add F-18
2. Evaporate
4. React
Cassette 2
Cassette 3
Cassette 1
3. Add MeCN
Read only
Validate operation parameters when React is clicked

## Slide 29
Manual Run (Run Operation)
Sequence
Abort
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Move
Prompt
1. Add F-18
2. Evaporate
4. React
Cassette 2
Cassette 3
Cassette 1
3. Add MeCN
Read only
Unit 1 Video
Reacting unit 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes
Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm

## Slide 30
Manual Run (Abort)
Sequence
Abort
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Move
Prompt
1. Add
2. Evaporate
4. React
Cassette 2
Cassette 3
Cassette 1
3. Add MeCN
Unit 1 Video
Reacting unit 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes
Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm
Are you sure you want to abort the current operation? This cannot be undone.
Abort
Cancel

## Slide 31
Manual Run (Error)
Sequence
Abort
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Move
Prompt
1. Add
2. Evaporate
4. React
Cassette 2
Cassette 3
Cassette 1
3. Add MeCN
Unit 1 Video
Reacting unit 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes
Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm
Fatal error: Hardware failure in unit 1: unable to close reactor.
OK

## Slide 32
Manual Run (Run Complete)
Sequence
Run Complete
Elute
Transfer
Install
Evaporate
React
Add
Toolbox
Addition of acetonitrile to unit 1 completed successfully.
Please select the next operation from the toolbox.
Move
Prompt
1. Add
2. Evaporate
3. React
Cassette 2
Cassette 3
Cassette 1
Click/tap to select the next unit operation
Read only (scrollable but not selectable)
3. Add
Specify a sequence name if you would like to add this run to your saved sequences.
Save
Cancel
Please enter a brief comment that describes the manual run:
FAC
FAC (short reaction times)

## Slide 33
This section describes the details for each unit operation:
View
Edit
Run

## Slide 34
Add Operation
View:
Add 10 mg tribenzoyl pentose triflate in 0.5 mL acetonitrile to reactor 1.
Edit:
Reactor:
1
Precursor1
Reagent:
Run:
Adding 10 mg tribenzoyl pentose triflate in
0.5 mL acetonitrile (MeCN) to reactor 1.
Temperature: 39.3 ⁰C
Activity: 15 mCi at 4:53 pm
Reactor 1 Video
This operation is only needed on the new Elixys system.

## Slide 35
Evaporate Operation (1)
View:
Evaporate reactor 1 at 105 ⁰C for 3 minutes and cool to 35 ⁰C. Stir at 500 rpm.
Edit:
Reactor:
1
3
Duration:
Evaporation temperature:
Final temperature:
Stir speed:
80.0
35.0
500
sec
⁰C
⁰C
rpm
0
min

## Slide 36
Evaporate Operation (2)
Run:
Reactor 1 Video
Evaporating reactor 1 at 105 ⁰C for 3 minutes and cooling to 35 ⁰C. Stirring at 500 rpm.
Status: Heating
Temperature: 78.9 ⁰C
Activity: 15 mCi at 4:53 pm

## Slide 37
Transfer Operation
View:
Transfer the contents of reactor 1 to the Silica gel cartridge.
Edit:
Run:
Reactor 1 Video
Reactor:
1
Target:
Column1
Transferring the contents of reactor 1 to the Silica gel cartridge.
Temperature: 39.3 ⁰C
Activity: 15 mCi at 4:53 pm

## Slide 38
Elute Operation
View:
Elute the Silica gel cartridge with ethyl acetate into reactor 2.
Edit:
Run:
Reactor 1 Video
Reactor:
1
Reagent:
MeCN1
Target:
Unit 2
Eluting the Silica gel cartridge with ethyl acetate into reactor 2.
Temperature: 39.3 ⁰C
Activity: 15 mCi at 4:53 pm

## Slide 39
React Operation (1)
View:
React reactor 1, position 1 at 160 ⁰C for 15 minutes and cool to 35 ⁰C.
Stir at 5000 rpm.
Edit:
Reactor:
1
15
Duration:
Reaction temperature:
Final temperature:
Stir speed:
160.0
35.0
5000
sec
⁰C
⁰C
rpm
0
min
Position:
1

## Slide 40
React Operation (2)
Run:
Reactor 1 Video
Reacting reactor 1, position 1 at 160 ⁰C for 15 minutes and cooling to 35 ⁰C. Stirring at 5000 rpm.
Status: Reacting, 8:23 minutes elapsed
Temperature: 112.0 ⁰C
Activity: 15 mCi at 4:53 pm

## Slide 41
Prompt Operation (1)
View:
Prompt the user "Please add 1 mL methanol/dichloromethane to
reactor 2 (1 of 5)"
Edit:
Message:
Please add 1 mL methanol/dichloromethane to reactor 2 (1 of 5)

## Slide 42
Prompt Operation (2)
Please add 1 mL methanol/
dichloromethane to reactor 2 (1 of 5)
OK
Abort

## Slide 43
Move Operation
View:
Move reactor 2 to position 1 and close.
Edit:
Run:
This operation is only needed on the original ARC-P system.
Reactor:
2
Position:
1
State:
Close
Moving reactor 2 to position 1 and closing.
Temperature: 39.3 ⁰C
Activity: 12 mCi at 4:53 pm

## Slide 44
Install Operation (1)
View:
Move reactor 2 to the install position and prompt the user “Take an
aliquot of the reaction for HPLC analysis.”
Edit:
Run:
This operation is only needed on the new Elixys system.
Reactor:
2
Moving reactor 2 to the install position.
Temperature: 39.3 ⁰C
Activity: 15 mCi at 4:53 pm
Message:
Take an aliquot of the reaction for HPLC analysis.

## Slide 45
Install Operation (2)
Take an aliquot of the reaction
for HPLC analysis.
OK
Abort
This operation is only needed on the new Elixys system.

## Slide 46
Comment Operation
View:
This is a comment entered by the user to make a note,
label a section of the synthesis reaction, etc.
Edit:
Comment:
This is a comment entered by the user to make a note, label a section of the synthesis reaction, etc.
Comment operations are ignored during the sequence run.

## Slide 47
Read Activity Operation
View:
Read the radiation activity level of the vial in reactor 2.
Edit:
Run:
This operation is only needed on the new Elixys system.
Reactor:
2
Reading the radiation activity level of the vial in reactor 2.
Temperature: 39.3 ⁰C
Activity: 15 mCi at 4:53 pm